## Slide 1
HiRA CsI Calibration for protons
(Low Energy DEE points + punch through)

## Slide 2
Data from runs:
2570 – 2573
2620 – 2630
2575 – 2578
2592 – 2611
DE-E plots for each individual CsI with strip front/CsI matching condition.
example from telescope 0 CsI 1
Zoom on
punch through points

## Slide 3
Fits for punch through points determination

## Slide 4
Calibration fits for telescope 0 (linear assumption)
Low energy points
from DE-E
punch-through
points

## Slide 5
Forward telescope
Backward telescope
DE-E with all CsI merged in the same plot
for each telescope

## Slide 6
DE-E line theoretically predicted for
1537 um Silicon
0.29 mg/cm2 Mylar (2.076 um)
10 cm CsI
The CsI energy at low channel values is under-estimate (shift of 5MeV for a Si energy of 11.3 MeV).
Calibration problems at low energy